Meta Data for North Pole Environmental Observatory 2002 Aerial CTD-O2 Survey

| Cast | - Station | Latitude N | Longitude E | Date | Time UTC | Soc | Voffset | sat depth m | “500” depth m |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Helo 0 | 88º32.4’ | 087º53.6’ | 4/23/02 | 10:14 | 0.6198 | 0.9027 | 53.108 | 500.183 |
| 2 | Helo 1 | 88º20.7’ | 113º03.0’ | 4/23/02 | 13:34 | 0.4894 | 0.4069 | 30.171 | 500.111 |
| 3 | Helo 2 | 88º19.4’ | 123º21.0’ | 4/23/02 | 14:32 | 0.5371 | 0.5575 | 4.081 | 500.169 |
| 4 | Helo 3 | 88º18.5’ | 131º59.0’ | 4/23/02 | 15:30 | 0.5052 | 0.4375 | 5.284 | 500.272 |
| 5 | Helo 4 | 88º16.3’ | 140º51.7’ | 4/23/02 | 16:26 | 0.4907 | 0.3837 | 8.403 | 500.131 |
| 6 | Helo 5 | 87º55.3’ | 153º50.9’ | 4/26/02 | 12:04 | 0.4144 | 0.1533 | 8.461 | 500.088 |
| 7 | Helo 6 | 88º06.2’ | 148º05.5’ | 4/26/02 | 12:59 | 0.4288 | 0.2154 | 10.828 | 500.132 |
| 8 | Helo 7 | 88º21.1’ | 118º23.7’ | 4/26/02 | 14:06 | 0.4761 | 0.4285 | 29.885 | 500.174 |
| 9 | Helo 8 | 88º30.1’ | 071º14.1’ | 4/28/02 | 10:41 | 0.4521 | 0.3268 | 30.049 | 500.178 |
| 10 | Hydro 1 | 85º10.7’ | 165º16.4’ | 4/27/02 | 14:33 | 0.4260 | 0.2021 | 12.776 | 500.075 |
| 11 | Hydro 2 | 86º37.1’ | 164º59.5’ | 4/27/02 | 18:50 | 0.4141 | 0.1417 | 8.075 | 500.114 |

- The data are available in the form of tab-delimited ASCII files with the following 10 column headers: scan, depth m, press db, temp C, cond S/m, sal, pot temp C, sigma theta, O2 V, O2 ml/L.  The name of file is the station name given in Table 1.  Temperatures are expressed on the ITS-90 scale and salinities as PSS78.  “O2 V” is the lag corrected oxygen voltage.

- Measurements were made with a Seabird SBE-19 Seacat CTD outfitted with an SEB43 O2-sensor as part of the observational program of NPEO’02, following landing of either a ML8 helicopter or Twin Otter aircraft on the sea ice.  Data from the SBE 43 dissolved O2 sensor were generated for all of the casts.  Individual dissolved sensor profiles generally appeared reasonable in shape, however, they exhibited hysteresis between the down and up casts and were all at least 1 ml/L less than the expected values. The O2 sensor profiles tended to converge at about 500 m on a given day but were offset between days by amounts (0.5-1.0 ml/L) much larger than likely environmental variability.  Obviously the sensor exhibited some type of problem.

- Our initial working hypothesis was that degassing of the electrolyte resulted in bubbles behind the membrane.  Such bubbles could cause divergence from the certified calibration parameters and hence offsets but not necessarily incorrect profile shapes.  While this was a common problem with early serial numbers of SBE43 sensors, it was determined that ours was a later serial number that should not have been subject to this problem.  Fouling or oil/grease contamination of the membrane can also diminish the sensitivity of the sensor but it seemed unlikely to us that this had occurred.  After the 2003 deployment, we learned that the O2 sensor contains a pressure compensation bladder and that Seabird had been filling this with distilled water through March 2003.  It is likely that during shipment and/or deployment the contents of the compensation bladder experienced a hard freeze and so stretched the sensor membrane such that its sensitivity was diminished but not eliminated.  Unfortunately, Niskin bottle O2 samples taken in 2002 were not trustworthy.  Thus the profiles were processed in a non-standard manner as follows.

CTD processing followed the SEASOFT recipe to minimize salinity spiking, with certain constants determined empirically.  The CTD-O2 package was set to acquire data at 2 scans/second.  Conductivity and temperature were low-pass filtered with a time constant of 0.3 seconds, pressure was filtered with a time constant of 2.0 seconds and temperature was advanced relative to pressure by 0.7 seconds.  Seasoft’s thermal lag correction was applied with alpha=0.03 and tau=9.0.  Potential temperature, depth and sigma-theta parameters were calculated according to EOS80 (Fofonoff and Millard, 1983).

- Typically, data from downcasts are preferred since the water column can be disturbed by the wire and instrument package during the upcast.  However, the upper portion of profiles in downcasts from the helicopter tended to be subject to ice interference in one or more of the sensors.  Ignoring this portion of the profile, the best alignment of the adjusted in-situ temperature and O2 voltages occurred for an O2 advance of 5.0 sec.  Comparison of processed up and downcasts shows that the O2 sensor signals tended to converge below 200 m but showed varying degrees of hysteresis above 200 m.  Accordingly, the following discussion is restricted to data from upcasts to 2 m below the surface for all sensors.  Above 2 m the data are subject to artifacts related to the ice hole.  The upcast data are available as ASCII files the title of which is the station identifier in the table above.  Also listed are parameters relevant to the calibration of the O2 data as follows.

Aligned oxygen voltage data were processed using a modified variant of the Owens and Millard algorithm (Owens and Millard, 1985) as described in the Sept 2002 Seabird Application Note No. 64:

O2 (ml/L) = {Soc * (V + Voffset)} * O2sat(T,S) * e (Tcor * T) * e(Pcor * P)

where Soc is the oxygen signal slope, V is the SBE43 temperature compensated output oxygen signal (Ox volts), Voffset is the voltage at zero oxygen signal, O2sat(T,S) is the oxygen saturation value in ml/L calculated using in-situ temperature and salinity following Weiss (1970), and Tcor and Pcor are residual temperature and pressure correction factors applied to in-situ temperature and pressure.  Slightly varying calibration parameters Soc, and Voffset for each profile are given in the table above. Values of 1.6x10-3 for Tcor, and 1.350x10-4 for Pcor based on the pre-mission calibration were applied to all of the profiles.  The other calibration parameters Soc, and Voffset were adjusted for each profile to give saturation in the surface mixed layer at the depth given in Table 1 and to bring concentrations at about 500m to 6.95 ml/L (Swift, et al., 1997).  The specific depth adjusted to 6.95 ml/L for each profile is also given in the table above.  Spikes in the O2 data were identified by visual inspection and eliminated though spiking was negligible except for portions of Helo-04 (cast 5).  We estimate that the data have an uncertainty of about 0.2 ml/L.  We have provided the aligned voltages, should anyone wish to undertake a different strategy in processing this data.  Obviously, the processed form of the O2 data should not be used to examine issues related to surface saturation state nor small differences at depth where all the profiles were forced to fit a given value.

This data was collected under the support of NSF grant OPP-9910335 to K. Falkner and OPP-9910305 to J. Morison.  Any opinions, findings, and conclusions or recommendations expressed in this material are those of the author(s) and do not necessarily reflect the views of the National Science Foundation.

For further information, please contact:

Dr. Kelly K. Falkner		kfalkner@coas.oregonstate.edu	(541) 737-3625

or Robert W. Collier		collier@coas.oregonstate.edu		(541) 737-4367

104 Ocean Admin Bldg

COAS, OSU

Corvallis, OR 97331-5503